## Sheet1
Inside diameter of ball	25	mm
diameter of opening	6	mm	recommended to be 20% of the ball outside diameter
density of liquid	1.00	kg/l

volume of ball	8.18E+03	mm2
upwards force	80.26E-03	N

surface area opening	28.27E+00	mm2

maximum pressure above	2.84E+03	Pa=N/m2
theoretical maximum level above opening	28.4	cm	(cm H2O = mBar = hPa)
advised minimum level above opening	5.7	cm
advised maximum level above opening	14.2	cm

If the water above the opening has the height of the theoretical maximum, the ball needs to be submerged completely to be able to close the hole. Any more and the level regulation mechanism cannot even work in theory anymore. Additionally, in practise extra force is required to compress the rubber seal to overcome imperfections. / The recommended maximum level is 50% of the theoretical maximum, to maintain closure pressure. The water level has then already risen 50% of the diameter of the ball. The minimum recommended level is 20%, in which case around 10% of the ball is submerged (see sheet2). The curve has sufficient gradient from 20% level onwards. / The opening cannot be made too small, because then the ball will not find the center by itself, and consequently the hole will not be closed.

## Sheet2
ball submerge %	ball volume %
0	0
5	0.725
10	2.8
15	6.075
20	10.4
25	15.625
30	21.6
35	28.175
40	35.2
45	42.525
50	50
55	57.475
60	64.8
65	71.825
70	78.4
75	84.375
80	89.6
85	93.925
90	97.2
95	99.275
100	100